## Slide 1: The GlucoWatch@ biographer: a frequent, automatic and noninvasive glucose monitor
Presenter: Zong Fan
Michael J. Tierney, Janet A. Tamada, Russell O. Potts, Richard C. Eastman, Kenneth Pitzer, Neil R. Ackerman and Steven J. Fermi, Cygnus Research Team
Ann Med. 2000 Dec;32:632-41. doi: 10.3109/07853890009002034.

## Slide 2: Agenda
Author introduction
Significance of research
Methods
Experiment results
Conclusion & Critique

## Slide 3: 1. Author introduction
Cygnus research team (Cygnus Inc., CA):
  1985-2007, not profitable
  Major business: develop and manufacture transfermal drug delivery system
  Launch GlucoWatch at 2002, first commercial non-invasive glucose monitoring product
Michael Tierney
  Chemistry Ph.D of Umich, currently in Google Research Department for sensor development
  14 papers about glucose monitoring sensor since 1999
Notes: products: contraception, homorn therapy, smoking cessation

## Slide 4: 2. Significance of research
Fingerstick blood glucose measurement is the gold standard but frequency is inefficient due to pain, inconvenience and cost.
Previous technologies to address the issue:
  Near infrared (NIR) spectroscopy
    No commercial product
    Low sensitivity of glucose signal
  Implantable subcutaneous sensor
    Poor assessibility of data
    Inconvenient calibration
GlucoWatch
  enable frequent and non-invasive measurement
  allow detection of patterns in glucose level for better glycaemic control
Notes: NIR spectra are made up of broad bands corresponding to overlapping peaks: the overtones (ie, first, second, third, and combination overtones), are formed by molecular vibrations.

## Slide 5: 3 Method - Structure
1. crystal display
2. electronics circuits
- amperometric biosensor
- iontophoresis
3. microprocessor
4. data memory & serial port
5. temperature & skin conductivity sensor
6. disposable AutoSensor
- biosensor
- iontophoresis electrodes
- hydrogel discs

## Slide 6: Principle
Figure 1. Schematic design of the GlucoWatch biographer
Working procedure:
1. glucose sample extraction via reverse iontophoresis
2. glucose measurement by amperometic sensor
3. data processing to display

## Slide 7: Principle (continue)
skin
blood vessel
electric current
Na+
glucose
Figure 2 Reverse iontophoresis to extract glucose
electro-omostic flow
Step 1
Step 2
Average flux: 1 nmol/cm2/L
Step 3
Figure 3 iontophoresis and biosensor duty cycle

## Slide 8: 4. Experiment Results
Table 1 summary of 6 clinical trials for evaluating the performance of GlucoWatch
Statistics of trial subjects:
1. gender: # of male ≈ # of female
2. age: 48
3. duration of diabetes: 17 years
4. ethnicity: 75% Caucasian, 10% Africa American, 15% others

## Slide 9: GlucoWatch could measure glucose accurately
Table 2 Accuracy measure of 4 trial studies
MD: mean differance
SD: standard deviation
MRD: mean of relative difference
MARD: mean of absolute value of relative difference
RMSE: root mean squared error

## Slide 10: Factors that affect measurement accuracy
Factor 1. blood glucose range
Table 3 Effect of blood glucose range on GlucoWatch accuracy

## Slide 11: Factors that affect GlucoWatch accuracy(continue)
Table 4 Effect of rate of change of blood glucose on GlucoWatch accuracy
Factor 3. Ethnicity
  African-Americans have negative values for MD
Factor 4. Body mass index
Factor 2. change rate of blood glucose level

## Slide 12: 5. Conclusion
Clinical trials show GlucoWatch biographer are accurate and precise compared to serial fingerstick glucose measurement over a 12-h period.
This noninvasive method extends the glucose measurement time range and provides accurate glycaemic alert.
Efficient data enables analysis of trend and pattern of glucose levels for better and personalized glycaemic control regimen.

## Slide 13: Critique
Higher Hyperglycarmic and Hypoglycaemic alert accuracy (75% and 79%, respectively) are required for practical use.
Biosafety is a big problem (0.8% strong skin irritation).
Practically poor performance, relatively complex operation, and imcompability with existing monitoring standard lead its failure.
Thinking:
  It's easier to get uncharged molecules into the skin than out of it.
  Trying to force glucose across the intact skin may alter the local concentration of glucose.
Notes: 75% with 10% false alerts
79% of hyperglycaemic points (BG > 17 mmol/L) with 8% false alerts.

## Slide 14: Reference
Tierney MJ, Tamada JA, Potts RO, Eastman RC, Pitzer K, Ackerman NR, Fermi SJ. The GlucoWatch biographer: a frequent automatic and noninvasive glucose monitor. Ann Med. 2000 Dec;32(9):632-41. doi: 10.3109/07853890009002034. PMID: 11209971.
Smith, John. (2006). The Pursuit of Noninvasive Glucose: Hunting the Deceitful Turkey.
So, Chi-Fuk et al. “Recent advances in noninvasive glucose monitoring.” Medical devices (Auckland, N.Z.) vol. 5 (2012): 45-52. doi:10.2147/MDER.S28134
Wikepedia - Cygnus Inc
Wikipedia - Noninvasive glucose monitor